## Studies
Boisson 2018 | Intra-operative cutaneous temperature monitoring with zero-heat-flux technique (3M SpotOn) in comparison with oesophageal and arterial temperature: A prospective observational study
Gómez-Romero 2019 | Intra-operative temperature monitoring with two non-invasive devices (3M Spoton® and Dräger Tcore®) in comparison with the Swan-Ganz catheter
Tachibana 2019 | Using the Bair Hugger™ temperature monitoring system in neck and chest regions: a pilot study
Schell-Chaple 2018 | Rectal and Bladder Temperatures vs Forehead Core Temperatures Measured With SpotOn Monitoring System.
Carvalho 2019 | Intra-operative Temperature Monitoring with Cutaneous Zero-Heat- Flux-Thermometry in Comparison with Oesophageal Temperature: a prospective study in paediatric population
KollmannCamaiora 2019 | Validation of the Zero-Heat-Flux thermometer (SpotOn) in major gynecological surgery to monitor intraoperative core temperature: a comparative study with esophageal core temperature
Idei 2016 | An evaluation and accuracy of new zero-heat-flux thermometer (3 M SpotOn) in pediatric intensive care patients
Iden 2015 | Intraoperative temperature monitoring with zero heat flux technology (3M SpotOn sensor) in comparison with sublingual and nasopharyngeal temperature
Kato 2015 | An evaluation of deep-forehead temperature (Spoton) in ICU patients after cardiac surgery
West 2018 | Agreement of a zero-heat-flux core temperature system with measurements obtained from nasal and oral temperature probes during general anesthesia
Pesonen 2018 | The focus of temperature monitoring with zero-heat-flux technology (3M Bair-Hugger): a clinical study with patients undergoing craniotomy
Makinen 2016 | Novel Zero-Heat-Flux Deep Body Temperature Measurement in Lower Extremity Vascular and Cardiac Surgery
Jack 2019 | Determining the accuracy of zero-flux and ingestible thermometers in the peri-operative setting
Eshraghi 2014 | An evaluation of a zero-heat-flux cutaneous thermometer in cardiac surgical patients
Dahyot-Fizelier 2017 | Accuracy of Zero-Heat-Flux Cutaneous Temperature in Intensive Care Adults

## Boisson 2018
Agreed
Study Identification
Sponsorship source | This study was supported by University Hospital of Poitiers and the SpotON temperature sensors and monitors were provided by 3M.
Country | France
Setting | perioperative
Comments
Authors name | Matthieu Boisson
Institution | Centre Hospitalier Universitaire de Poitiers
Email | matthieu.boisson@chu-poitiers.fr
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Adult patients anesthetized for elective major abdominal surgery were enrolled.
Exclusion criteria | Patients with any contra-indication to oesophageal probe insertion or undergoing emergency surgery were excluded.
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Median age (year) | 55 (IQR 47-69) | 64 (IQR 55-69)
Female (n) | 9 | 9
Male (n) | 13 | 33
Temperature range (Celsius) | 33.0 to 49.8 | 34.0 to 40.2
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
| Slow core temperature changes |  |  |  |  | Rapid core temperature changes
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.1 | -0.5 | 0.3 | 1708 | 22 | 0.5 | -1.2 | 2.2 | 72 | 22
SpotON minus oesphageal | 0.1 | -0.4 | 0.6 | 2345 | 49 | 0.6 | -1.2 | 2.4 | 101 | 49
Dan Phan
Study Identification
Sponsorship source | This study was supported by University Hospital of Poitiers and the SpotON temperature sensors and monitors were provided by 3M.
Country | France
Setting | perioperative
Comments
Authors name | Matthieu Boisson
Institution | Centre Hospitalier Universitaire de Poitiers
Email | matthieu.boisson@chu-poitiers.fr
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Adult patients anesthetized for elective major abdominal surgery were enrolled.
Exclusion criteria | Patients with any contra-indication to oesophageal probe insertion or undergoing emergency surgery were excluded.
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Median age (year) | 55 |  | 64
Female (n) | 16 | 9
Male (n) | 33 | 13
Temperature range (Celsius) | 34.0 to 40.2 | 33.0 to 39.8
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
| Slow core temperature changes |  |  |  |  | Rapid core temperature changes
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.1 | -0.5 | 0.3 | 1708 | 22 | 0.5 | -1.2 | 2.2 | 72 | 22
SpotON minus oesphageal | 0.1 | -0.4 | 0.6 | 2345 | 27 | 0.6 | -1.2 | 2.4 | 101 | 27
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Median age (year)
Female (n)
Male (n)
Temperature range (Celsius)
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
| Slow core temperature changes |  |  |  |  | Rapid core temperature changes
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.1 | -0.5 | 0.3 | 1708 | 22 | 0.5 | -1.2 | 2.2 | 72 | 22
SpotON minus oesphageal | 0.1 | -0.4 | 0.6 | 2345 | 49 | 0.6 | -1.2 | 2.4 | 101 | 49

## Gómez-Romero 2019
Agreed
Study Identification
Sponsorship source
Country | Spain
Setting | Intraoperative
Comments
Authors name | Gomez-Romero
Institution | Research department, Ctudad Real University General Hospital
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients undergoing programmed valvular surgery with sufficient surface area on the forehead to accommodate both devices (3M SpotOn and Drager Tcore) simultaneously as well as placement of the Swanz-Ganze catheter were included.
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years) |  | 69 ± 10.8
Female (n) |  | 58
Male (n) |  | 231
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0.21 | -2.27 | 2.71 | 289 | 41
Dan Phan
Study Identification
Sponsorship source
Country | Spain
Setting | Intraoperative
Comments
Authors name | Gomez-Romero
Institution | Research department, Ctudad Real University General Hospital
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients undergoing programmed valvular surgery with sufficient surface area on the forehead to accommodate both devices (3M SpotOn and Drager Tcore) simultaneously as well as placement of the Swanz-Ganze catheter were included.
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years) |  | 69
Female (n) |  | 58
Male (n) |  | 231
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0.21 | -2.27 | 2.71 | 289 | 41
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years)
Female (n)
Male (n)
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0.21 | -2.27 | 2.71 | 289 | 41

## Tachibana 2019
Agreed
Study Identification
Sponsorship source
Country | Japan
Setting | Intraoperative
Comments
Authors name | Shunsuke tachibana
Institution | Department of Anesthesiology, Sapporo Medical University
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Female patients between 21 and 80 years old who had scheduled laparoscopic surgery under general anesthesia.
Exclusion criteria | Patients with esophageal lesions and/or abnormalities in the region of attachment and expected to be suffering from massive hemorrhage during surgery were excluded.
Group differences
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal | Overall
Mean age (years) | 49.3 ± 15.9 | 43.4 ± 14.7 | 58.6 ± 17.9
Female (n) | 6 | 5 | 6
Male (n) | 4 | 5 | 4
Interventions
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger head placement minus esophageal | 0.01 | -0.47 | 0.49 | 303 | 10
Bair Hugger neck placement minus esophageal | 0.05 | -0.3 | 0.4 | 446 | 10
Bair Hugger chest placement minus esophageal | -0.55 | -1.1 | 0 | 487 | 10
Dan Phan
Study Identification
Sponsorship source
Country | Japan
Setting | Intraoperative
Comments
Authors name | Shunsuke tachibana
Institution | Department of Anesthesiology, Sapporo Medical University
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Female patients between 21 and 80 years old who had scheduled laparoscopic surgery under general anesthesia.
Exclusion criteria | Patients with esophageal lesions and/or abnormalities in the region of attachment and expected to be suffering from massive hemorrhage during surgery were excluded.
Group differences
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal | Overall
Mean age (years) | 49.3 | 43.4 | 59.6
Female (n) | 6 | 5 | 6
Male (n) | 4 | 5 | 4
Interventions
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger head placement minus esophageal | 0.01 | -0.47 | 0.49 | 303 | 10
Bair Hugger neck placement minus esophageal | 0.05 | -0.3 | 0.5 | 446 | 10
Bair Hugger chest placement minus esophageal | -0.55 | -1.1 | 0 | 487 | 10
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal | Overall
Mean age (years)
Female (n)
Male (n)
Interventions
| Bair Hugger head placement minus esophageal | Bair Hugger neck placement minus esophageal | Bair Hugger chest placement minus esophageal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger head placement minus esophageal | 0.01 | -0.47 | 0.49 | 303 | 10
Bair Hugger neck placement minus esophageal | 0.05 | -0.3 | 0.4 | 446 | 10
Bair Hugger chest placement minus esophageal | -0.55 | -1.1 | 0 | 487 | 10

## Schell-Chaple 2018
Agreed
Study Identification
Sponsorship source
Country | United States of America
Setting | Intensive Care Unit
Comments
Authors name | Hildy M. Schell-Chaple
Institution | University of California, San Francisco Medical Center
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients enrolled in a randomized clinical trial to test the effects of intravenous acetaminophen on core temperature and hemodynamic responses in febrile critically ill patients were eligible for the study. Of these, patients were eligible for enrollment if they were 18 years or older, weighed at least 50 kg (110 lb) and had a fever (temperature ≥ 38.3°C)
Exclusion criteria | Patients with malignant hyperthermia or heat stroke, acute liver failure, therapeutic temperature management or extra-corporeal blood circuit therapy were excluded.
Group differences
| SpotON minus bladder | SpotON minus rectal | Overall
Mean age (years) |  |  | 57 ± 15
Female (n) |  |  | 18
Male (n) |  |  | 20
Temperature range SpotON forehead | 36.7 to 39.9 | 36.6 to 39.4
Interventions
| SpotON minus bladder | SpotON minus rectal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus bladder | -0.07 | -0.54 | 0.4 | 180 | 9
SpotON minus rectal | -0.24 | -0.81 | 0.33 | 568 | 29
Dan Phan
Study Identification
Sponsorship source
Country | United States of America
Setting | Intensive Care Unit
Comments
Authors name | Hildy M. Schell-Chaple
Institution | University of California, San Francisco Medical Center
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients enrolled in a randomized clinical trial to test the effects of intravenous acetaminophen on core temperature and hemodynamic responses in febrile critically ill patients were eligible for the study. Of these, patients were eligible for enrollment if they were 18 years or older, weighed at least 50 kg (110 lb) and had a fever (temperature ≥ 38.3°C)
Exclusion criteria | Patients with malignant hyperthermia or heat stroke, acute liver failure, therapeutic temperature management or extra-corporeal blood circuit therapy were excluded.
Group differences
| SpotON minus bladder | SpotON minus rectal | Overall
Mean age (years) |  |  | 57
Female (n) |  |  | 18
Male (n) |  |  | 20
Temperature range SpotON forehead | 36.7 to 39.9 | 36.6 to 39.4
Interventions
| SpotON minus bladder | SpotON minus rectal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus bladder | -0.07 | -0.54 | 0.4 | 180 | 9
SpotON minus rectal | -0.24 | -0.81 | 0.33 | 568 | 29
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus bladder | SpotON minus rectal | Overall
Mean age (years)
Female (n)
Male (n)
Temperature range SpotON forehead
Interventions
| SpotON minus bladder | SpotON minus rectal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus bladder | -0.07 | -0.54 | 0.4 |  | 9
SpotON minus rectal | -0.24 | -0.81 | 0.33 |  | 29

## Carvalho 2019
Agreed
Study Identification
Sponsorship source | Universitair Ziekenhuis Brussel
Country | Belgium
Setting | Intraoperative
Comments
Authors name | Hugo Carvalho
Institution | Department of Anesthesiology and Perioperative Medicine, Universitair Ziekenhuis Brussel.
Email | carvalho.hn@gmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Children aged 1-12 years with American Society of Anaesthesiology physical status classification 1 or 2 scheduled to undergo elective surgical procedures under general anesthesia for a minimum of 30 minutes.
Exclusion criteria | Fragile forehead skin state; known allergy to the probe adhesive or any constituent components; maxillofacial trauma or lesions; procedures impeding proper placement of the SpotOn™ sensor; abnormal esophageal anatomy/gastro‐esophageal procedures; coagulopathy; neurologically impaired children with abnormal thermoregulation; hemodynamic instability; need for vasoactive medication; procedures associated with extended use of abdomino/thoracic rinsing fluids; thoracoscopic/thoracotomy procedures; malignant hyperthermia or family history of malignant hyperthermia; patient with fever or infection; all conditions that might be judged to abnormally alter skin perfusion; anatomical variants (overt hydrocephalus …); and calibration/device failure.
Group differences
| SpotON minus esophageal | Overall
Mean Age (months) |  | 66.1 ± 46.8
Female (n) |  | 12
Male (n) |  | 42
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.14 | -0.39 | 0.66 | 1664 | 54
Dan Phan
Study Identification
Sponsorship source | Universitair Ziekenhuis Brussel
Country | Belgium
Setting | Intraoperative
Comments
Authors name | Hugo Carvalho
Institution | Department of Anesthesiology and Perioperative Medicine, Universitair Ziekenhuis Brussel.
Email | carvalho.hn@gmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Children aged 1-12 years with American Society of Anaesthesiology physical status classification 1 or 2 scheduled to undergo elective surgical procedures under general anesthesia for a minimum of 30 minutes.
Exclusion criteria | Fragile forehead skin state; known allergy to the probe adhesive or any constituent components; maxillofacial trauma or lesions; procedures impeding proper placement of the SpotOn™ sensor; abnormal esophageal anatomy/gastro‐esophageal procedures; coagulopathy; neurologically impaired children with abnormal thermoregulation; hemodynamic instability; need for vasoactive medication; procedures associated with extended use of abdomino/thoracic rinsing fluids; thoracoscopic/thoracotomy procedures; malignant hyperthermia or family history of malignant hyperthermia; patient with fever or infection; all conditions that might be judged to abnormally alter skin perfusion; anatomical variants (overt hydrocephalus …); and calibration/device failure
Group differences
| SpotON minus esophageal | Overall
Mean Age (months) |  | 66.1 ± 46.8
Female (n) |  | 12
Male (n) |  | 42
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.14 | -0.39 | 0.66 | 1664 | 54
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus esophageal | Overall
Mean Age (months)
Female (n)
Male (n)
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.14 | -0.39 | 0.66 | 1664 | 54

## KollmannCamaiora 2019
Agreed
Study Identification
Sponsorship source
Country | Spain
Setting | Intraoperative
Comments
Authors name | Andrea Kollmann Camaiora
Institution | Department of Anesthesia and Intensive Care, La Paz University Hospital
Email | and_c@hotmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients undergoing gynecological surgery between May and December 2016. Further inclusion criteria include a BMI between 20-35 kg/m2, ages 18-80 years, and an ASA physical status of 1-3.
Exclusion criteria | Exclusion criteria were patients with a febrile process in the last seven days, hypothermia (<34) prior to surgery, medical contraindications for oral intubation, medical contraindications for the use of sevoflurane, a risk of blood loss of >1000mL, and thyroid dysfunction.
Group differences
| SpotON minus esophageal | Overall
Mean age (years) |  | 51.5 ± 15.2
Female (n) |  | 70
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Overall
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | -0.27 | -0.61 | 0.55 | 401 | 66
Dan Phan
Study Identification
Sponsorship source
Country | Spain
Setting | Intraoperative
Comments
Authors name | Andrea Kollmann Camaiora
Institution | Department of Anesthesia and Intensive Care, La Paz University Hospital
Email | and_c@hotmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients undergoing gynecological surgery between May and December 2016. Further inclusion criteria include a BMI between 20-35 kg/m2, ages 18-80 years, and an ASA physical status of 1-3.
Exclusion criteria | Exclusion criteria were patients with a febrile process in the last seven days, hypothermia (<34) prior to surgery, medical contraindications for oral intubation, medical contraindications for the use of sevoflurane, a risk of blood loss of >1000mL, and thyroid dysfunction.
Group differences
| SpotON minus esophageal | Overall
Mean age (years) |  | 51.5
Female (n) |  | 70
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Overall
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | -0.27 | -0.61 | 0.55 | 401 | 66
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus esophageal | Overall
Mean age (years)
Female (n)
Interventions
SpotON minus esophageal
Outcomes
Mean Bias
Overall
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | -0.27 | -0.61 | 0.55 | 401 | 66

## Idei 2016
Agreed
Study Identification
Sponsorship source
Country | Japan
Setting | Pediatric Intensive Care Unit
Comments
Authors name | M. Idei
Institution | Yokohoma City University Hospital, Intensive Care Unit
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Pediatric patients weighing less than 10kg who were staying in the ICU during the period from February 2015 to March 2016
Exclusion criteria
Group differences
| SpotON minus rectal | Overall
Patients (n) |  | 26
Interventions
SpotON minus rectal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus rectal | 0.82 | 0.31 | 1.33 | 53495 | 26
Dan Phan
Study Identification
Sponsorship source
Country | Japan
Setting | Pediatric Intensive Care Unit
Comments
Authors name | M. Idei
Institution | Yokohoma City University Hospital, Intensive Care Unit
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Pediatric patients weighing less than 10kg who were staying in the ICU during the period from February 2015 to March 2016
Exclusion criteria
Group differences
| SpotON minus rectal | Overall
Patients (n) |  | 26
Interventions
SpotON minus rectal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus rectal | 0.82 | 0.31 | 1.33 | 53495 | 26
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus rectal | Overall
Patients (n)
Interventions
SpotON minus rectal
Outcomes
Mean bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus rectal | 0.82 | 0.31 | 1.33 | 53495 | 26

## Iden 2015
Agreed
Study Identification
Sponsorship source | SpotON temperature sensors were provided by 3M
Country | Germany
Setting | Intraoperative
Comments
Authors name | Timo Iden
Institution | Department of Anaesthesiology and Intensive Care Medicine University Medical Center Schleswig-Holstein
Email | timo.iden@uksh.de
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Men and women undergoing elective trauma or gynaecology surgery under general anesthesia were eligible
Exclusion criteria | Patients with unexpected blood loss, haemodynamic instability determined by the need of continuous norepinephrine infusion or needing post-operative ventilation were excluded
Group differences
| SpotON minus nasopharyngeal | SpotON minus sublingual | Overall
Mean age female (years) |  |  | 47.7 ± 14.1
Mean age male (years) |  |  | 55.0 ± 16.8
Female (n) |  |  | 55
Male (n) |  |  | 28
Interventions
| SpotON minus nasopharyngeal | SpotON minus sublingual
Outcomes
Mean Bias
| 15 min |  |  |  |  | 45 min |  |  |  |  | 75 min
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasopharyngeal | 0.07 | -0.38 | 0.51 | 83 | 83 | 0.05 | -0.39 | 0.48 | 83 | 83 | 0.1 | -0.25 | 0.46 | 83 | 83
SpotON minus sublingual | -0.37 | -0.95 | 0.22 | 83 | 83 | -0.36 | -0.95 | 0.23 | 83 | 83 | -0.33 | -0.84 | 0.19 | 83 | 83
Dan Phan
Study Identification
Sponsorship source | SpotON temperature sensors were provided by 3M
Country | Germany
Setting | Intraoperative
Comments
Authors name | Timo Iden
Institution | Department of Anaesthesiology and Intensive Care Medicine University Medical Center Schleswig-Holstein
Email | timo.iden@uksh.de
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Men and women undergoing elective trauma or gynaecology surgery under general anesthesia were eligible
Exclusion criteria | Patients with unexpected blood loss, haemodynamic instability determined by the need of continuous norepinephrine infusion or needing post-operative ventilation were excluded
Group differences
| SpotON minus nasopharyngeal | SpotON minus sublingual | Overall
Mean age female (years) |  |  | 47.7
Mean age male (years) |  |  | 55
Female (n) |  |  | 55
Male (n) |  |  | 28
Interventions
| SpotON minus nasopharyngeal | SpotON minus sublingual
Outcomes
Mean Bias
| 15 min |  |  |  |  | 45 min |  |  |  |  | 75 min
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasopharyngeal | 0.07 | -0.38 | 0.51 | 83 | 83 | 0.05 | -0.39 | 0.48 | 83 | 83 | 0.1 | -0.25 | 0.46 | 83 | 83
SpotON minus sublingual | -0.37 | -0.95 | 0.22 | 83 | 83 | -0.36 | -0.95 | 0.23 | 83 | 83 | -0.33 | -0.84 | 0.19 | 83 | 83
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus nasopharyngeal | SpotON minus sublingual | Overall
Mean age female (years)
Mean age male (years)
Female (n)
Male (n)
Interventions
| SpotON minus nasopharyngeal | SpotON minus sublingual
Outcomes
Mean Bias
| 15 min |  |  |  |  | 45 min |  |  |  |  | 75 min
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasopharyngeal | 0.7 | -0.38 | 0.51 | 83 | 83 | 0.05 | -0.39 | 0.48 | 83 | 83 | 0.1 | -0.25 | 0.46 | 83 | 83
SpotON minus sublingual | -0.37 | -0.95 | 0.22 | 83 | 83 | -0.36 | -0.95 | 0.23 | 83 | 83 | -0.33 | -0.84 | 0.19 | 83 | 83

## Kato 2015
Agreed
Study Identification
Sponsorship source
Country | Japan
Setting | Postoperative
Comments
Authors name | H. Kato
Institution | Hamamatsu University School of Medicine, Department of Anesthesia and Intensive Care, Hamamatsu, Japan
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Postoperative cardiac surgery patients were included
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years) |  | 66 (range 45-78 years)
Female (n) |  | 4
Male (n) |  | 16
Temperature range of pulmonary artery (Celsius) |  | 35.6 to 39.2
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.28 | -1.16 | 0.6 | 16407 | 20
Dan Phan
Study Identification
Sponsorship source
Country | Japan
Setting | Postoperative
Comments
Authors name | H. Kato
Institution | Hamamatsu University School of Medicine, Department of Anesthesia and Intensive Care, Hamamatsu, Japan
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Postoperative cardiac surgery patients were included
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years) |  | 66
Female (n) |  | 4
Male (n) |  | 16
Temperature range of pulmonary artery (Celsius) |  | 35.6 to 39.2
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.28 | 1.16 | 0.6 | 16407 | 20
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus arterial | Overall
Mean age (years)
Female (n)
Male (n)
Temperature range of pulmonary artery (Celsius)
Interventions
SpotON minus arterial
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | -0.28 | -1.16 | 0.6 | 16407 | 20

## West 2018
Agreed
Study Identification
Sponsorship source | 3M Canada
Country | Canada
Setting | Intraoperative
Comments | 3M Canada sponsored the trial and provided the Bair Hugger core temperature monitoring system and sensors. Dan Mores from 3M provided the statistical analysis.
Authors name | Nicholas West
Institution | Digital Health Innovation Lab, BC Children's Hospital Research Institute and Department of Anesthesiology, Pharmacology and therapeutics, University of British Columbia
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | ASA I-III patients scheduled from elective, non-cardiac surgery under general anesthesia
Exclusion criteria
Group differences
| SpotON minus nasal | SpotON minus oral | Overall
Number of patients | 170` | 24 | 194
Median age (years) | 60 (IQR 48-68) | 62 (IQR 54-72)
Female (n) | 80 | 15 | 95
Male (n) | 90 | 9 | 99
Interventions
| SpotON minus nasal | SpotON minus oral
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasal | -0.11 | -0.75 | 0.53 | 109570 | 170
SpotON minus oral | -0.06 | -0.64 | 0.52 | 22482 | 24
Dan Phan
Study Identification
Sponsorship source | 3M Canada
Country | Canada
Setting | Intraoperative
Comments | 3M Canada sponsored the trial and provided the Bair Hugger core temperature monitoring system and sensors. Dan Mores from 3M provided the statistical analysis.
Authors name | Nicholas West
Institution | Digital Health Innovation Lab, BC Children's Hospital Research Institute and Department of Anesthesiology, Pharmacology and therapeutics, University of British Columbia
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | ASA I-III patients scheduled from elective, non-cardiac surgery under general anesthesia
Exclusion criteria
Group differences
| SpotON minus nasal | SpotON minus oral | Overall
Number of patients | 170` | 24 | 194
Median age (years) | 60 | 62
Female (n) | 80 | 15 | 95
Male (n) | 90 | 9 | 99
Interventions
| SpotON minus nasal | SpotON minus oral
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasal | -0.11 | -0.75 | 0.53 | 109570 | 170
SpotON minus oral | -0.06 | -0.64 | 0.52 | 22482 | 24
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus nasal | SpotON minus oral | Overall
Number of patients
Median age (years)
Female (n)
Male (n)
Interventions
| SpotON minus nasal | SpotON minus oral
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus nasal | -0.11 | -0.75 | 0.53 | 109570 | 170
SpotON minus oral | -0.06 | -0.64 | 0.52 | 22482 | 24

## Pesonen 2018
Agreed
Study Identification
Sponsorship source | University of Helsinki including Helsinki University Central Hospital. The study was supported by a grant from the Viipuri Tuberculosis Foundation and by Hospital University EVO grants
Country | Finland
Setting | Intraoperative
Comments
Authors name | Eero Pesonen
Institution | Division of Anesthesiology, Department of Anesthesiology, Intensive Care and Pain Medicine, University of Helsinki and Helsinki University Hospital
Email | eero.pesonen@hus.fi
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Adults patients scheduled for elective craniotomy
Exclusion criteria
Group differences
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder | Overall
Median age (years) |  |  | 59 (IQR 44-67)
Male (n) |  |  | 10
Female (n) |  |  | 19
Interventions
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair-Hugger minus naso-pharyngeal | 0.11 | -0.54 | 0.75 |  | 29
Bair-Hugger minus bladder | -0.14 | -0.81 | 0.52 |  | 29
Dan Phan
Study Identification
Sponsorship source | University of Helsinki including Helsinki University Central Hospital. The study was supported by a grant from the Viipuri Tuberculosis Foundation and by Hospital University EVO grants
Country | Finland
Setting | Intraoperative
Comments
Authors name | Eero Pesonen
Institution | Division of Anesthesiology, Department of Anesthesiology, Intensive Care and Pain Medicine, University of Helsinki and Helsinki University Hospital
Email | eero.pesonen@hus.fi
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Adults patients scheduled for elective craniotomy
Exclusion criteria
Group differences
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder | Overall
Median age (years) |  |  | 59
Male (n) |  |  | 10
Female (n) |  |  | 19
Interventions
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair-Hugger minus naso-pharyngeal | 0.11 | -0.54 | 0.75 |  | 29
Bair-Hugger minus bladder | -0.14 | -0.81 | 0.52 |  | 29
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments | Conflict of interest: Author M-TM has been on a member of 3M patient temperature management advisory board until the end of 2016. No other conflicts of interest to declare
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder | Overall
Median age (years)
Male (n)
Female (n)
Interventions
| Bair-Hugger minus naso-pharyngeal | Bair-Hugger minus bladder
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair-Hugger minus naso-pharyngeal | 0.11 | -0.54 | 0.75 |  | 29
Bair-Hugger minus bladder | -0.14 | -0.81 | 0.52 |  | 29

## Makinen 2016
Agreed
Study Identification
Sponsorship source | SpotON temperature monitoring system was supplied for by 3M
Country | Finland
Setting | Intraoperative
Comments | The author Marja-Tellervo Makinen is a member of an advisory board of 3M for patient temperature management. This study was investigator initiated.
Authors name | Marja-Tellervo Makinen
Institution | Meilahti Hospital, University of Helsinki and Helsinki University Hospital
Email | makinentelle@gmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients scheduled for vascular surgery of the lower extremities or undergoing cardiac surgery with cardiopulmonary bypass surgery for cornary artery bypass grafting, aortic valve replacement or reconstruction of the ascending aorta were included.
Exclusion criteria
Group differences
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal | Overall
Mean age (years) |  |  |  | 64 ± 13
Female (n) |  |  |  | 9
Male (n) |  |  |  | 21
Interventions
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal
Outcomes
Mean Bias
| Vascular Surgery |  |  |  |  | Cardiac Surgery
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.08 | -0.25 | 0.4 |  | 15
SpotON minus arterial | [5 empty] | -0.05 | -0.56 | 0.47 |  | 15
SpotON minus nasopharyngeal | [5 empty] | -0.12 | -0.94 | 0.71 |  | 15
Dan Phan
Study Identification
Sponsorship source | SpotON temperature monitoring system was supplied for by 3M
Country | Finland
Setting | Intraoperative
Comments | The author Marja-Tellervo Makinen is a member of an advisory board of 3M for patient temperature management. This study was investigator initiated.
Authors name | Marja-Tellervo Makinen
Institution | Meilahti Hospital, University of Helsinki and Helsinki University Hospital
Email | makinentelle@gmail.com
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients scheduled for vascular surgery of the lower extremities or undergoing cardiac surgery with cardiopulmonary bypass surgery for cornary artery bypass grafting, aortic valve replacement or reconstruction of the ascending aorta were included.
Exclusion criteria
Group differences
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal | Overall
Mean age (years) |  |  |  | 64
Female (n) |  |  |  | 9
Male (n) |  |  |  | 21
Interventions
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal
Outcomes
Mean Bias
| Vascular Surgery |  |  |  |  | Cardiac Surgery
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.08 | -0.25 | 0.4 |  | 15
SpotON minus arterial | [5 empty] | 0.05 | -0.56 | 0.47 |  | 15
SpotON minus nasopharyngeal | [5 empty] | -0.12 | -0.94 | 0.71 |  | 15
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal | Overall
Mean age (years)
Female (n)
Male (n)
Interventions
| SpotON minus esophageal | SpotON minus arterial | SpotON minus nasopharyngeal
Outcomes
Mean Bias
| Vascular Surgery |  |  |  |  | Cardiac Surgery
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus esophageal | 0.08 | -0.25 | 0.4 |  | 15
SpotON minus arterial | [5 empty] | -0.05 | -0.56 | 0.47 |  | 15
SpotON minus nasopharyngeal | [5 empty] | -0.12 | -0.94 | 0.71 |  | 15

## Jack 2019
Agreed
Study Identification
Sponsorship source | This study was funded by a small project grant from the Association of Anesthetists of Great Britain and Ireland Foundation, administered by the National Institute for Academic Anesthesia
Country | Great Britain and Ireland
Setting | Perioperative
Comments | Author C. Mark Harper has received loans of equipment from variousmanufacturers of warming devices, including Inditherm, Augustine Biomedical, Arizant, 3M and Mölnlycke. He has also received expenses and honoraria for sitting on advisory boards for 3M and Mölnlycke
Authors name | James Jack
Institution | Brighton and sussex University Hospitals Trust, Royal Sussex County Hospital
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients eligible for inclusions were those over 18 years of age and listed for elective surgery at the authors' hospital trust either at the Royal Sussex County Hospital, Brighton or Princess royal Hospital, Haywards Heath
Exclusion criteria | Exclusion criteria included the inability to give informed consent, contraindication to insertion of oesophageal probe (carcinoma of the oesophagus or pharynx, previous oesophageal surgery, oesophageal stricture or varices, pharyngeal pouch), subjects undergoing open intra-abdominal surgery to remove a portion of the gastrointestinal tract, and abnormal gastrointestinal pathology that may prevent safe use of ITS.
Group differences
| Bair Hugger minus esophageal | Overall
Median age (years) |  | 48 (IQR 42-68)
Female (n) |  | 22
Male (n) |  | 7
Interventions
Bair Hugger minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger minus esophageal | 0.02 | -0.5 | 0.5 | 2511 | 29
Dan Phan
Study Identification
Sponsorship source | This study was funded by a small project grant from the Association of Anesthetists of Great Britain and Ireland Foundation, administered by the National Institute for Academic Anesthesia
Country | Great Britain and Ireland
Setting | Perioperative
Comments | Author C. Mark Harper has received loans of equipment from variousmanufacturers of warming devices, including Inditherm, Augustine Biomedical, Arizant, 3M and Mölnlycke. He has also received expenses and honoraria for sitting on advisory boards for 3M and Mölnlycke
Authors name | James Jack
Institution | Brighton and sussex University Hospitals Trust, Royal Sussex County Hospital
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients eligible for inclusions were those over 18 years of age and listed for elective surgery at the authors' hospital trust either at the Royal Sussex County Hospital, Brighton or Princess royal Hospital, Haywards Heath
Exclusion criteria | Exclusion criteria included the inability to give informed consent, contraindication to insertion of oesophageal probe (carcinoma of the oesophagus or pharynx, previous oesophageal surgery, oesophageal stricture or varices, pharyngeal pouch), subjects undergoing open intra-abdominal surgery to remove a portion of the gastrointestinal tract, and abnormal gastrointestinal pathology that may prevent safe use of ITS.
Group differences
| Bair Hugger minus esophageal | Overall
Median age (years) |  | 48
Female (n) |  | 22
Male (n) |  | 7
Interventions
Bair Hugger minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger minus esophageal | 0.02 | -0.5 | 0.5 | 2511 | 29
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| Bair Hugger minus esophageal | Overall
Median age (years)
Female (n)
Male (n)
Interventions
Bair Hugger minus esophageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
Bair Hugger minus esophageal | 0.02 | -0.05 | 0.05 | 2511 | 29

## Eshraghi 2014
Agreed
Study Identification
Sponsorship source | Funded by 3M
Country | United States of America
Setting | Intraoperative and postoperatively
Comments | Dr. Sessler serves on two 3M advisory boards and Dr. Sessler's department conducts sponsored trials for 3M and other companies that manufacture temperature monitoring and measurement systems
Authors name | Yashar Eshraghi
Institution | Department of Anesthesiology, Case Western Reserve University, Metrohealth Medical Center, Cleveland, Ohio
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients having elective cardiac surgery with cardiopulmonary bypass and a pulmonary artery catheter.
Exclusion criteria | Patients were excluded when deep hypothermic cardiac arrest was planned as were patients with forehead or neck skin rash or infection.
Group differences
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery | Overall
Mean age (years) |  |  | 67 ± 9
Temperature ranges (Celsius) |  |  | 38
Male (n) |  |  | 67
Temperature range for pulmonary artery (Celsius) |  |  | 32.0 to 39.3
Interventions
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery
Outcomes
Mean Bias
| Overall |  |  |  |  | Intraoperative |  |  |  |  | Postoperative
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
Forehead SpotON placement minus pulmonary artery | -0.23 | -1.05 | 0.59 | 35717 | 103 | -0.08 | -0.96 | 0.8 | 13929 | 103 | -0.32 | -1.07 | 0.43 | 21788 | 103
Lateral neck SpotON placement minus pulmonary artery | -0.3 | -1.18 | 0.58 | 35717 | 103 | -0.15 | -0.99 | 0.69 | 13929 | 103 | -0.4 | -1.24 | 0.44 | 21788 | 103
Dan Phan
Study Identification
Sponsorship source | Funded by 3M
Country | United States of America
Setting | Intraoperative and postoperatively
Comments | Dr. Sessler serves on two 3M advisory boards and Dr. Sessler's department conducts sponsored trials for 3M and other companies that manufacture temperature monitoring and measurement systems
Authors name | Yashar Eshraghi
Institution | Department of Anesthesiology, Case Western Reserve University, Metrohealth Medical Center, Cleveland, Ohio
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | Patients having elective cardiac surgery with cardiopulmonary bypass and a pulmonary artery catheter.
Exclusion criteria | Patients were excluded when deep hypothermic cardiac arrest was planned as were patients with forehead or neck skin rash or infection.
Group differences
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery | Overall
Mean age (years) |  |  | 67
Temperature ranges (Celsius) |  |  | 38
Male (n) |  |  | 67
Temperature range for pulmonary artery (Celsius) |  |  | 32.0 to 39.3
Interventions
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery
Outcomes
Mean Bias
| Overall |  |  |  |  | Intraoperative |  |  |  |  | Postoperative
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
Forehead SpotON placement minus pulmonary artery | -0.23 | -1.05 | 0.65 | 35717 | 103 | -0.08 | -0.96 | 0.8 | 13929 | 103 | -0.32 | -1.07 | 0.43 | 21788 | 103
Lateral neck SpotON placement minus pulmonary artery | -0.3 | -1.18 | 0.58 | 35717 | 103 | -0.15 | -0.99 | 0.69 | 13929 | 103 | -0.4 | -1.24 | 0.44 | 21788 | 103
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery | Overall
Mean age (years)
Temperature ranges (Celsius)
Male (n)
Temperature range for pulmonary artery (Celsius)
Interventions
| Forehead SpotON placement minus pulmonary artery | Lateral neck SpotON placement minus pulmonary artery
Outcomes
Mean Bias
| Overall |  |  |  |  | Intraoperative |  |  |  |  | Postoperative
| mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants | mean bias | lower LoA | upper LoA | measurements | participants
Forehead SpotON placement minus pulmonary artery | -0.23 | -1.05 | 0.59 | 35717 | 103 | -0.08 | -0.96 | 0.8 | 13929 | 103 | -0.32 | -1.07 | 0.43 | 21788 | 103
Lateral neck SpotON placement minus pulmonary artery | -0.3 | -1.18 | 0.58 | 35717 | 103 | -0.15 | -0.99 | 0.69 | 13929 | 103 | -0.4 | -1.24 | 0.44 | 21788 | 103

## Dahyot-Fizelier 2017
Agreed
Study Identification
Sponsorship source | Funded by University Hospital of Poitiers
Country | France
Setting | Surgical and Neuro Intensive Care Units
Comments | Author Dr. Mimoz received funding from 3M
Authors name | Claire Dahyot-Fizelier
Institution | University Hospital of Poitiers
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | All adult patients admitted in the neurosurgical or surgical ICU from August to November 2014 were eligible for inclusion into this study.
Exclusion criteria | Exclusion criteria were skin injury or surgical suture in the forehead and esophageal probe contraindication.
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Mean age (years) |  |  | 54 ± 5
Femae (n) |  |  | 25
Male (n) |  |  | 27
Temperature range (Celsius) |  |  | 33.0 to 39.7
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0 | -0.36 | 0.36 | 1850 | 7
SpotON minus oesphageal | -0.19 | -0.72 | 0.34 | 61298 | 52
Dan Phan
Study Identification
Sponsorship source | Funded by University Hospital of Poitiers; Author Dr. Mimoz received funding from 3M
Country | France
Setting | Surgical and Neuro Intensive Care Units
Comments
Authors name | Claire Dahyot-Fizelier
Institution | University Hospital of Poitiers
Email
Address
Methods
Design | Prospective cohort study
Group
Population
Inclusion criteria | All adult patients admitted in the neurosurgical or surgical ICU from August to November 2014 were eligible for inclusion into this study.
Exclusion criteria | Exclusion criteria were skin injury or surgical suture in the forehead and esophageal probe contraindication.
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Mean age (years) |  |  | 54
Femae (n) |  |  | 25
Male (n) |  |  | 27
Temperature range (Celsius) |  |  | 33.0 to 39.7
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0 | -0.36 | 0.36 | 1850 | 7
SpotON minus oesphageal | -0.19 | -0.72 | 0.34 | 61298 | 52
Megan Bittner
Study Identification
Sponsorship source
Country
Setting
Comments
Authors name
Institution
Email
Address
Methods
Design
Group
Population
Inclusion criteria
Exclusion criteria
Group differences
| SpotON minus arterial | SpotON minus oesphageal | Overall
Mean age (years)
Femae (n)
Male (n)
Temperature range (Celsius)
Interventions
| SpotON minus arterial | SpotON minus oesphageal
Outcomes
Mean Bias
Baseline
| mean bias | lower LoA | upper LoA | measurements | participants
SpotON minus arterial | 0 | -0.36 | 0.36 | 1850 | 7
SpotON minus oesphageal | -0.19 | -0.72 | 0.34 | 61298 | 52